## STX
Robotic magnetic navigation system with accessories for complex electrophysiology and catheter ablation procedures
| Subject of procurement | quantity (pcs)
| Complete equipment of the electrophysiological hall with a robotic magnetic navigation system with accessories, including a 3D electroanatomical mapping system, electrophysiological recording system, external stimulator, radio frequency generator and flushing pump. | 3
Technical specification
| Technical properties | Required item
1. | Robotic magnetic navigation system for electrophysiology | yes
| Robotic magnetic system with complete hardware and software equipment directly dedicated to electrophysiology and catheter ablation, including complex forgings | yes
| Ability to remotely control and navigate RMN (Remote Magnetic Navigation) catheters in both endocardial and epicardial anatomy of the heart with a high level of accuracy and safety and maintained high flexibility of controlled catheters | yes
| Control the procedure from the control room (mouse, joystick or keyboard) | yes
| Ability to control procedures remotely, from anywhere, using a computer network | yes
| Saving optional "one physician setup" settings for the ablation procedure to improve workflow and optimize and minimize human resources | yes
| Graphical integration and display of catheter position, mapping software and X-ray imaging to minimize X-ray exposure and maximize procedure success | yes
| Minimization of X-rays for the patient and the doctor | yes
| Semi-automatic RMN catheter navigation to improve workflow | yes
| Possibility to use automated functions, such as automatic mapping, moving the magnetic catheter to a specified position or moving the magnetic catheter along a specified trajectory using two-way communication with a 3D mapping system | yes
| Maximum possibility of integration and cooperation between the robotic magnetic navigation system and the 3D mapping system, including the transmission of information about the contact of the magnetic catheter with the tissue and the parameters of the ablation lesion | yes
| Possibility to use robotic magnetic navigation system for procedures with manual control of standard ablation and diagnostic catheters without robotic magnetic navigation | yes
Monitors
| 2 pieces of QHD 58 "monitors (1 pc in the operating room and 1 pc in the control room) with integrated and compatible display of video sources from the 3D electroanatomical mapping system and electrophysiological system without reducing the image quality | yes
| Integrated simultaneous display of min. 11 external video sources, including: X-ray, 3D electroanatomical mapping system, electrophysiological system, intracardiac ultrasound, hemodynamic monitoring and others. | yes
| Adjustable location and size of all application windows, including windows displaying connected third-party devices | yes
| Change and save window layout settings for individual users | yes
| Integrated control of all connectable devices (3D mapping system, RMN, RTG) on one terminal and monitor | yes
| Create and export large monitor desktop images and video files (screenshot) | yes
| Output of a digital video signal for the transmission of the image of a large-area monitor to the lecture room via an IP network | yes
| Open software interface for possible integration of additional mapping systems | yes
| Noise filter for Biosense Webster Carto mapping system | yes
| An operating room monitor with a ceiling mount that allows the monitor to be positioned in all directions, at least in range: left / right, down / up, closer / farther to ensure a good view of the monitor by the operator and allow it to be moved up or shifted longitudinally to free up space for manipulating the patient during movements in the operating room | yes
Digital image processing system - post-processing
| Acquisition matrix > 1548x1548px | yes
| Processing depth > 14 bit | yes
| Pulse sciascopy with a variable frequency in the range of at least 3.75 to 15 fps. | yes
| Image memory on HDD X-ray system without additional consoles, discs and external storage devices as well as CD / DVD ≥ 80,000 images | yes
| Possibility to import information from a 3D scan or 3D map and overlay it with an X-ray image in the integrated robotic navigation system | yes
| Zoom in, panning, flipping in post-processing | yes
| Digital system control panel in the control room | yes
| Implementation of digital system functions from the control panel in the control room | yes
| Image storage and recording on CD / DVD in the available DICOM 3.0 standard (Send, Inquiry / Get, Receive, Saving workbooks with browsing layering) | yes
Operating table
| Operating table length | min. 260 cm
| Operating table top width | min. 45 cm
| Longitudinal movement of the operating table | min. 120 cm
| Maximum load capacity of the patient's operating table | min. 200 kg
| Lateral movement of the operating table top | min. +/- 12 cm
| Patient table absorbance 1.0 mm Al | yes
| Mattress and pad for the patient's head | yes
| Integrated infusion bottle stand | yes
| Cable holder at the table | yes
Imaging (registration) system
| Ability to hold the last X-ray image (LIH) in the display | yes
| DICOM 3.0 interface (minimum Save, Send, Print, Inquiry / Received, Commitment, Workbook) | yes
| Saving images to CD / DVD / R / RW storage in DICOM 3.0 standard | yes
Control console
| Operators' workspace is equipped with a 58-inch QHD monitor to control the robotic navigation system and integrated third-party devices | yes
| Control panel for controlling the X-ray movement of the C-arm in the control console | yes
| Control interface for controlling the movement of the catheter tip | yes
| Emergency system switch | yes
| Accessories | yes
| Operating LED light installed on an adjustable articulated ceiling hinge | yes
| Medical tripod with articulated ceiling suspension with a load capacity of min. 300 kg, variably adjustable position and height, equipped with structured cabling and adapted to electrophysiological devices and accessories | yes
| Communication equipment, including a camera system with the possibility of video transmission to the lecture room | yes
2. | 3D electroanatomical mapping system
| Electromagnetic sensor contact 3D mapping system including RF generator with flush pump and remote control | yes
| Possibility of full integration with a robotic magnetic navigation system using a magnetic ablation catheter and magnetic localization | yes
| Possibility to control the navigation of the magnetic ablation catheter directly in the application of the 3D anatomical mapping system | yes
| Possibility of creating electro-anatomical maps of individual anatomical structures of the heart and extracardiac structures (especially atria, ventricles, thoracic veins, aorta) using an ablation navigation catheter or navigation diagnostic catheter | yes
| Possibility of creating anatomical maps and reconstruction of individual anatomical structures of the heart and extracardiac structures (especially atria, ventricles, thoracic veins, aorta) using an intracardiac echo catheter | yes
| Possibility to determine the size of individual heart compartments | yes
| Possibility of displaying the time course of electrical activation, voltage information (unipolar and bipolar) or correlation of stimulation mapping of individual parts of cardiac compartments using color display | yes
| Possibility of detailed mapping of individual structures ("point by point"), fast anatomical mapping and automatic acquisition of points with automatic annotation using a catheter-heart contact filter using special multipolar diagnostic catheters and an ablation catheter | yes
| Ability to accurately locate and display ablation and multiple electrophysiological catheters in 3D reconstructed cardiac compartments in real time during exercise | yes
| Possibility of different way of displaying the magnitude of force / pressure and direction of action of the catheter on the tissue (graphically, numerically), using a fully integrated measurement system | yes
| Possibility of automatic acquisition of ablation points depending on power, pressure, stability and time | yes
| Possibility of integration of CT or MRI image into real 3D anatomy | yes
| Fast automatic mapping and sorting of ventricular extrasystoles and ventricular tachycardias | yes
| Possibility to use a controllable loader visible in the 3D map in real time | yes
| Support for catheters that allow the creation of maps with a high density of trigger points | yes
| Basic physical and technical requirements of the system: | yes
| System accuracy: less than 1mm in vivo | yes
| Sensor position detection in several directions and planes (min. 5) | yes
| Possibility to protect the system against shock from defibrillators | yes
3. | Electrophysiological recording system and stimulator
| Electrophysiological recording system designed for use in electrophysiological procedures with the possibility of monitoring and recording hemodynamic parameters and optional display of images on any monitors in the operating and control room. | yes
| Compatibility with 3D mapping system and robotic magnetic system with the possibility of integrated control. | yes
| Database with the possibility of storing results such as: patient demographic data associated with the time course of ECG waves, Intracardiac ECG, pressure and other measured parameters, as well as calculated indicators. | yes
| Computer with at least two color TFT / LCD monitors with a diagonal of min. 20 ” | yes
| Transmission of images of the electrophysiological recording system to the monitor of the digital angiography system LDM (Large Display Monitor) | yes
| Ergonomic set of amplifier connection modules (connection box) providing connection of at least 100 bipolar signals (minimum 200 individual inputs) | yes
Amplifier of electrical signals of the heart with at least:
| 100 intrakardiálnymi kanálmi (minimálne 200 jednotlivých vstupov) | yes
| 12 ECG channels | yes
| 4 pressure channels | yes
| 4 inputs for external rhythm stimulator | yes
| Communication between the amplifier and the control computer via an optical cable | yes
| Frame rate: 1, 2 and 4 [kHz] | yes
| Filter for frequency 50 [Hz] and 60 [Hz], designed and adjustable for individual intracardiac channels | yes
System requirements:
| Receiving signals | yes
| Control of ablation parameters | yes
| Storage of acquired data and their analysis | yes
| Browsing the acquired data | yes
| Real-time ECG display of intracardiac electrograms on the monitor | yes
| Display of recorded ECG signals and intracardiac electrograms and their analysis on a “post processing” browser | yes
| Interactive window - LOG enabling saving and editing of all events during the examination, as well as the possibility of complete editing of events and their entry | yes
| Measurements and presentation of current (current) patient rhythm cycles from any two selected channels | yes
| Possibility to set and change the ECG speed on the live monitor (real time) in the minimum range of 6 - 400 [mm / s], and on the viewing monitor in the range up to 1,600 [mm / s] | yes
| Number of channels simultaneously displayed on the monitor at least 64 in real time | yes
| Automatic data collection of ablation parameters and their presentation in digital graphic form | yes
| Automatic data collection of cryoablation parameters and their presentation in digital graphic form | yes
The system must allow the user to freely set channels, including the independent setting of signal parameters for each intracardiac channel:
| magnification / amplification | yes
| color | yes
| frequencies for filter groups | yes
| network filters | yes
| Possibility to select any channel for direct stimulation from the live (real-time) monitor without the need to intervene via the system menu | yes
| Ability to define the individual bipolar channels used as a stimulation channel and create an on-screen shortcut for them | yes
| Possibility to create preferential settings of pages with preset composition and graphic location of individual channels. | yes
| Possibility to set study protocols and assign individual events and intervals | yes
| Possibility of free specification of a sample window with signals (both 12 ECG and intracardiac signals) for performing and comparing morphology with subsequence recorded signals implemented in real time | yes
User interface with functions for setting and visualization of electrophysiological signals with definition:
| regulation of intensity, ECG signal strength, intracardiac (IK) signal, invasive pressure | yes
| horizontal and vertical adjustment of ECG signal pattern, IK, invasive pressure | yes
| Automatic detection of RF applications with automatic recording of stimulation parameters and their presentation in the window - LOG for a subsequence of ablation steps with the most used RF generators available on the market | yes
| Automatic recording of pacing pulses with automatic acquisition of pacing parameters and their presentation in the window - LOG for further pacing steps | yes
| Ability to create comments for the collected data and their simple future analysis | yes
| Possibility of analysis of ECG signals as well as intracardiac recordings | yes
| Manual and automatic measurement of intervals with analysis of ECG curves and intracardiac channels | yes
| Trigger and trigger mode (with the possibility of refreshing the screen) | yes
| Presentation and recording with the Holter system (Holter window) | yes
| Possibility to print ECG curves and intracardiac signals from the screen | yes
| Possibility of creating your own procedural reports | yes
| Automatic recording of total ablation time and cryoablation time information in reports | yes
| Save browsing screens and images in BMP and JPG format | yes
| Creating presentations in MS Power Point software, including exported data in Excel | yes
| Ability to export "RAW data" in binary or xls format | yes
| Software integration of electrophysiological recording system with electro-anatomical 3D mapping system | yes
| Emergency pacing-function to bypass the amplifier (directly from the stimulator) | yes
| Possibility of later modernization of the device for measuring hemodynamic parameters such as calculation of cardiac output, invasive and non-invasive pressure on one platform | yes
Other parameters
| Possibility of Barcode scanner used for registration of consumables used during procedures | yes
| Archiving of measured data on DVD or SD card, or on a network disk at the level of electrophysiological software | yes
| Fast storage of patient data on disk, including patient identification data | yes
| HL7 data export for hospital system | yes
| DICOM MWL import patient data via communication user interface | yes
| Sharing monitors, workstation and keyboard in case of expansion by hemodynamic module | yes
| UPS for electrophysiological station | yes
| Programmable stimulator | yes
| Programmable stimulator with at least two channels, based on a computer platform with its own keyboard, LCD touch monitor and control stimulator located in the control room | yes
| Programmable stimulator with programmable parameters: | yes
| amplitude of the stimulation pulse in the range of at least 0.5 - 20 [mA] adjustable in 0.1 [mA] steps | yes
| pulse length 0.5 [ms] and in the range 1 - 2.9 [ms] with adjustable rise 1 [ms] | yes
| Burst stimulation with a pulse length of at least 30 - 3000 [ms] | yes
4. | Other requirements
| New, unrefined, unused goods | yes
| Free service during the warranty period of 24 months | yes
| Commissioning of the device | yes
| Training min. 3 employees at the place of delivery | yes
| Service available within 24 hours | yes
| Transport to the place of performance | yes
| Assembly and installation of equipment into operation | yes
| Submission of documents necessary for its collection and use (delivery note, or acceptance protocol, invoice, operating manual in Slovak or Czech language) | yes
| Testing the functionality and operability of the delivered equipment | yes
| Educational program for employees | yes
,